## Particle standard curve
Number of Particles | 300000000 | 150000000 | 75000000 | 37500000 | 18750000 | 9375000 | 4687500 | 2343750 | 1171875 | 585937.5 | 292968.75 | 0
Replicate 1 | 1.164 | 0.53 | 0.308 | 0.149 | 0.81 | 0.92 | 0.52 | 0.42 | 0.46 | 0.42 | 0.38 | 0.4 | Enter Abs600 measurements into blue cells
Replicate 2 | 0.952 | 0.553 | 0.275 | 0.151 | 0.81 | 0.58 | 0.5 | 0.54 | 0.4 | 0.53 | 0.37 | 0.37 | Gold cells are calculated
Replicate 3 | 0.956 | 0.391 | 0.326 | 0.167 | 0.8 | 0.58 | 0.46 | 0.42 | 0.54 | 0.39 | 0.41 | 0.4
Replicate 4 | 1.052 | 0.43 | 0.231 | 0.129 | 0.957 | 0.63 | 0.5 | 0.42 | 0.45 | 0.4 | 0.39 | 0.4
Arith. Mean | 1.031 | 0.476 | 0.285 | 0.149 | 0.84425 | 0.6775 | 0.495 | 0.45 | 0.4625 | 0.435 | 0.3875 | 0.3925
Arith. Std.Dev. | 0.099993 | 0.077859 | 0.041737 | 0.015578 | 0.075314 | 0.16338 | 0.025166 | 0.06 | 0.057951 | 0.06455 | 0.017078 | 0.015
Arith. Net Mean | 0.6385 | 0.0835 | -0.1075 | -0.2435 | 0.45175 | 0.285 | 0.1025 | 0.0575 | 0.07 | 0.0425 | -0.005
Values should form a straight line on both linear and log scale
Slope should be 1:1
Common problems:
* Consistent pipetting error --> log graph is a straight line but not 1:1 slope
* Oversaturated detector --> low concentrations linear, but high concentrations saturate or fall
Cospheric Monodisperse Silica Microspheres 0.961um diameter
https://www.cospheric.com/SiO2MS_monodisperse_silica_spheres_beads_nm_microns.htm
[15 empty] | Spheres/gram | 1200000000000 | Number supplied by manufacturer
[15 empty] | grams/mL | 2 | Number supplied by manufacturer
[15 empty] | Spheres/0.50 mL | 1200000000000 | Shipped volume from manufacturer
[15 empty] | Resuspend volume mL: | 40 | Preparation by iGEM HQ
[15 empty] | Tube Particles / mL: | 30000000000 | Particle density in measurement kit tube
Particles / Abs600 | [14 empty] | Dilution for stock solution: | 10 | 100 uL from tube + 900 ddH20
Number of Particles | 300000000 | 150000000 | 75000000 | 37500000 | 18750000 | 9375000 | 4687500 | 2343750 | 1171875 | 585937.5 | 292968.75 |  |  |  | Stock Particles / mL: | 3000000000 | Particle density in prepared stock solution
Mean particles / Abs600 | 469851213.7823 | 1796407185.6287 | -697674418.6047 | -154004106.7762 | 41505257.3326 | 32894736.8421 | 45731707.3171 | 40760869.5652 | 16741071.4286 | 13786764.7059 | -58593750 |  |  |  | Well volume (mL) | 0.1 | Amount of stock left in initial well after protocol
Mean of med-high levels: |  | 203825730.8845 | [12 empty] | Initial particles / well: | 300000000 | Amount of particles left in first well
Scaling is computed from medium-high points likely to be less impacted by saturation or pipetting error
If needed, you can shift which points are used, but it is likely better to correct instrument settings and protocol.

## Fluorescein standard curve
Fluorescein uM | 10 | 5 | 2.5 | 1.25 | 0.625 | 0.3125 | 0.15625 | 0.078125 | 0.039062 | 0.019531 | 0.0097656 | 0
Replicate 1 | 122928 | 65295 | 31123 | 16068 | 8241 | 4326 | 2191 | 985 | 600 | 335 | 201 | 43 | Enter fluorescence measurements into blue cells
Replicate 2 | 131616 | 61824 | 32245 | 15485 | 8760 | 4411 | 2271 | 1170 | 612 | 327 | 184 | 34 | Gold cells are calculated
Replicate 3 | 126914 | 58760 | 30422 | 17613 | 7925 | 4124 | 2093 | 1012 | 587 | 303 | 162 | 37
Replicate 4 | 118013 | 57072 | 28322 | 15901 | 8252 | 4307 | 2169 | 1006 | 549 | 296 | 168 | 30 | Values measured are fluorescence from 100uL of X uM fluorescein solution
Arith. Mean | 124867.75 | 60737.75 | 30528 | 16266.75 | 8294.5 | 4292 | 2181 | 1043.25 | 587 | 315.25 | 178.75 | 36
Arith. Std.Dev. | 5787.2344 | 3619.2856 | 1651.2405 | 930.3757 | 345.393 | 120.7836 | 73.2302 | 85.2892 | 27.313 | 18.6971 | 17.5 | 5.4772
Arith. Net Mean | 124831.75 | 60701.75 | 30492 | 16230.75 | 8258.5 | 4256 | 2145 | 1007.25 | 551 | 279.25 | 142.75
Values should form a straight line on both linear and log scale
Slope should be 1:1
Common problems:
* Consistent pipetting error --> log graph is a straight line but not 1:1 slope
* Oversaturated detector --> low concentrations linear, but high concentrations saturate or fall
Fluorescein uM --> MEFL calculation:
[15 empty] | Initial Molarity | 1e-05 | 10 micro-molar, per protocol
[15 empty] | Molecules / Mole | 6.0221e+23 | Avogadro's Number
[15 empty] | Well volume (L): | 0.0001 | 100 micro-liters
[15 empty] | Initial Molecules / well: | 602214090000000 | molecules/mole * moles/liter * liters
[15 empty] | MEFL / uM | 60221409000000 | initial well molecules / 10 micromoles
Fluorescein/a.u.
Fluorescein uM | 10 | 5 | 2.5 | 1.25 | 0.625 | 0.3125 | 0.15625 | 0.078125 | 0.039062 | 0.019531 | 0.0097656
uM Fluorescein/a.u. | 8.0108e-05 | 8.237e-05 | 8.1989e-05 | 7.7014e-05 | 7.568e-05 | 7.3426e-05 | 7.2844e-05 | 7.7563e-05 | 7.0894e-05 | 6.9942e-05 | 6.8411e-05
Mean uM fluorescein / a.u.: |  | 7.8096e-05
MEFL / a.u.: |  | 4703031049.0016
Final scaling level determined from medium-high points likely to be less impacted by saturation or pipetting error
If needed, you can shift which points are used, but it is likely better to correct instrument settings and protocol.

## Raw Plate Reader Measurements
Raw Plate Readings |  |  |  | Copy Abs600 and fluorescence measurements from your plate reader into the blue and green cells of the plate rectangles
They will automatically propagate into the correct locations in the Fluorescence per Particle Sheet
Media Blank Control Wells: |  |  | Identify the plate wells containing your media blank controls in the blue cells; the ranges will be calculated in gold and highlighted in green eblow
First well row (A-H): |  |  | A |  | Abs600 blanks: |  | J12:J19
First well column (1-12): |  |  | 9 |  | Fluorescence blanks: |  | J24:J31
Last well row (A-H): |  |  | H |  | Abs600 blank mean: |  | 0.37875
Last well column (1-12): |  |  | 9 |  | Fluorescence blank mean: |  | 318.25
Abs600 Raw Readings:
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.71 | 0.697 | 0.325 | 0.63 | 0.639 | 0.669 | 0.215 | 0.594 | 0.38
B | 0.744 | 0.702 | 0.392 | 0.694 | 0.668 | 0.594 | 0.188 | 0.59 | 0.39
C | 0.728 | 0.723 | 0.348 | 0.642 | 0.571 | 0.591 | 0.179 | 0.544 | 0.36
D | 0.709 | 0.706 | 0.344 | 0.601 | 0.592 | 0.54 | 0.176 | 0.533 | 0.37
E | 0.688 | 0.667 | 0.618 | 0.644 | 0.656 | 0.591 | 0.188 | 0.668 | 0.38
F | 0.685 | 0.635 | 0.651 | 0.652 | 0.596 | 0.56 | 0.181 | 0.562 | 0.36
G | 0.669 | 0.633 | 0.609 | 0.679 | 0.568 | 0.544 | 0.184 | 0.595 | 0.37
H | 0.679 | 0.659 | 0.621 | 0.65 | 0.582 | 0.516 | 0.181 | 0.563 | 0.42
Fluorescence Raw Readings:
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 355 | 1917 | 6344 | 2670 | 458 | 13337 | 2531 | 1654 | 294
B | 340 | 1965 | 5394 | 2322 | 345 | 10233 | 2162 | 1545 | 300
C | 338 | 1964 | 5180 | 2490 | 355 | 9304 | 2172 | 1489 | 367
D | 352 | 1943 | 5297 | 2506 | 338 | 8879 | 2053 | 1554 | 310
E | 355 | 3200 | 1594 | 2749 | 404 | 11606 | 2291 | 2037 | 314
F | 353 | 1973 | 1383 | 2662 | 373 | 8297 | 1909 | 1292 | 313
G | 348 | 1933 | 1360 | 2609 | 362 | 8072 | 1875 | 1283 | 324
H | 340 | 1963 | 1388 | 2691 | 358 | 8943 | 1907 | 1289 | 324

## Fluorescence per Particle
Unit Scaling Factors: |  |  | Calculated values imported from prior sheets |  |  |  | Enter fluorescence and Abs600 measurements into blue cells on "Raw Plate Reader Measurements"
Particles / Abs600 |  |  | 203825730.8845 |  |  |  | Gold cells are calculated from values on other sheets
MEFL / a.u. |  |  | 4703031049.0016
Abs600 blank mean: |  |  | 0.37875
Fluorescence blank mean: |  |  | 318.25
Calibration ready? |  |  | TRUE
Control blanks: |  |  | J12:J19 | The MEFL/particle values for these cells will generally be invalid, and are highlighed in red.
Experimental Values:
MEFL / particle
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 2559.884 | 115912.6895 | -2586732.2548 | 215975.2174 | 12390.2456 | 1034941.7252 | -311795.5255 | 143186.1 | -447631.4249 | --- | --- | ---
B | 1374.0036 | 117546.0328 | 8839000.9921 | 146658.5132 | 2133.8764 | 1062814.4378 | -223026.4257 | 133991.7893 | -37430.8064 | --- | --- | ---
C | 1304.8167 | 110308.4425 | -3648096.6894 | 190353.2464 | 4410.7235 | 976844.578 | -214132.8088 | 163471.3074 | -59991.8405 | --- | --- | ---
D | 2358.0325 | 114558.0805 | -3305859.1661 | 227130.1358 | 2136.9625 | 1224985.4458 | -197421.6925 | 184852.0555 | 21755.2828 | --- | --- | ---
E | 2741.9938 | 230677.812 | 123036.075 | 211448.0768 | 7136.4366 | 1227095.9447 | -238630.7153 | 137106.5432 | -78450.8683 | --- | --- | ---
F | 2618.1682 | 149000.3722 | 90239.8985 | 197911.0086 | 5814.9124 | 1015723.9194 | -185611.2422 | 122608.9928 | 6460.6597 | --- | --- | ---
G | 2365.0133 | 146542.3559 | 104395.7234 | 176040.8743 | 5334.0982 | 1082652.7008 | -184442.1796 | 102938.4225 | -15162.7729 | --- | --- | ---
H | 1671.4565 | 135416.976 | 101891.3572 | 201837.135 | 4512.5852 | 1449949.9116 | -185377.8791 | 121567.8512 | 3216.3458 | --- | --- | ---
Equivalent particle count:
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 67517273.3555 | 64867538.854 | -10955633.035 | 51211214.8847 | 53045646.4627 | 59160418.3892 | -33376463.4323 | 43873488.5729 | 254782.1636 | --- | --- | ---
B | 74447348.2056 | 65886667.5084 | 2700690.9342 | 64256061.6613 | 58956592.6583 | 43873488.5729 | -38879758.1662 | 43058185.6494 | 2293039.4725 | --- | --- | ---
C | 71186136.5114 | 70167007.857 | -6267641.2247 | 53657123.6554 | 39185496.7625 | 43262011.3802 | -40714189.7442 | 33682202.0287 | -3821732.4541 | --- | --- | ---
D | 67313447.6246 | 66701970.432 | -7082944.1482 | 45300268.6891 | 43465837.1111 | 32866899.1051 | -41325666.9368 | 31440118.9889 | -1783475.1452 | --- | --- | ---
E | 63033107.276 | 58752766.9275 | 48765306.1141 | 54064775.1171 | 56510683.8877 | 43262011.3802 | -38879758.1662 | 58956592.6583 | 254782.1636 | --- | --- | ---
F | 62421630.0834 | 52230343.5392 | 55491555.2333 | 55695380.9642 | 44281140.0347 | 36943413.7228 | -40306538.2824 | 37351065.1846 | -3821732.4541 | --- | --- | ---
G | 59160418.3892 | 51822692.0774 | 46930874.5362 | 61198675.6981 | 38574019.5699 | 33682202.0287 | -39695061.0898 | 44077314.3038 | -1783475.1452 | --- | --- | ---
H | 61198675.6981 | 57122161.0804 | 49376783.3068 | 55287729.5024 | 41427579.8023 | 27975081.5639 | -40306538.2824 | 37554890.9155 | 8407811.399 | --- | --- | ---
Net Abs 600
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.33125 | 0.31825 | -0.05375 | 0.25125 | 0.26025 | 0.29025 | -0.16375 | 0.21525 | 0.00125 | --- | --- | ---
B | 0.36525 | 0.32325 | 0.01325 | 0.31525 | 0.28925 | 0.21525 | -0.19075 | 0.21125 | 0.01125 | --- | --- | ---
C | 0.34925 | 0.34425 | -0.03075 | 0.26325 | 0.19225 | 0.21225 | -0.19975 | 0.16525 | -0.01875 | --- | --- | ---
D | 0.33025 | 0.32725 | -0.03475 | 0.22225 | 0.21325 | 0.16125 | -0.20275 | 0.15425 | -0.00875 | --- | --- | ---
E | 0.30925 | 0.28825 | 0.23925 | 0.26525 | 0.27725 | 0.21225 | -0.19075 | 0.28925 | 0.00125 | --- | --- | ---
F | 0.30625 | 0.25625 | 0.27225 | 0.27325 | 0.21725 | 0.18125 | -0.19775 | 0.18325 | -0.01875 | --- | --- | ---
G | 0.29025 | 0.25425 | 0.23025 | 0.30025 | 0.18925 | 0.16525 | -0.19475 | 0.21625 | -0.00875 | --- | --- | ---
H | 0.30025 | 0.28025 | 0.24225 | 0.27125 | 0.20325 | 0.13725 | -0.19775 | 0.18425 | 0.04125 | --- | --- | ---
Net Fluorescein a.u.
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 36.75 | 1598.75 | 6025.75 | 2351.75 | 139.75 | 13018.75 | 2212.75 | 1335.75 | -24.25 | --- | --- | ---
B | 21.75 | 1646.75 | 5075.75 | 2003.75 | 26.75 | 9914.75 | 1843.75 | 1226.75 | -18.25 | --- | --- | ---
C | 19.75 | 1645.75 | 4861.75 | 2171.75 | 36.75 | 8985.75 | 1853.75 | 1170.75 | 48.75 | --- | --- | ---
D | 33.75 | 1624.75 | 4978.75 | 2187.75 | 19.75 | 8560.75 | 1734.75 | 1235.75 | -8.25 | --- | --- | ---
E | 36.75 | 2881.75 | 1275.75 | 2430.75 | 85.75 | 11287.75 | 1972.75 | 1718.75 | -4.25 | --- | --- | ---
F | 34.75 | 1654.75 | 1064.75 | 2343.75 | 54.75 | 7978.75 | 1590.75 | 973.75 | -5.25 | --- | --- | ---
G | 29.75 | 1614.75 | 1041.75 | 2290.75 | 43.75 | 7753.75 | 1556.75 | 964.75 | 5.75 | --- | --- | ---
H | 21.75 | 1644.75 | 1069.75 | 2372.75 | 39.75 | 8624.75 | 1588.75 | 970.75 | 5.75 | --- | --- | ---